## Sheet1
Term | Definition | Example usage | Bioinformatics translation | Bioinformatics example sentence | Ontological reference(s) | Link to reference
Acceptability criteria | Predefined specifications or limits used to determine whether a result, process, or product is considered fit for purpose. Must be documented before testing, e.g., control values must fall within a certain range, PCR efficiency must meet specified thresholds. | Acceptability criteria for the qPCR assay required that positive control Ct values fall within ±1.0 of the established mean and that the negative control showed no amplification before results were approved. | Predefined quality control (QC) or performance thresholds applied to sequencing data and pipelines. Examples include minimum read quality (e.g., Phred ≥ 30), acceptable mapping/alignment rate (e.g., ≥ 90%), and validated benchmark performance (e.g., sensitivity, specificity, and PPV for variant calling). May also cover reproducibility and runtime expectations for bioinformatics workflows. | Acceptability criteria required that at least 90% of exome target bases achieved a minimum coverage of 20×
Accreditation | Formal recognition by an authoritative body that a laboratory or organisation is competent to perform specific tests or activities in accordance with established standards (e.g., ISO 15189). | The laboratory received ISO 15189 accreditation for molecular diagnostics, confirming it met all requirements for quality management, personnel competence, and technical performance. |  |  | mesh:D000068 [Certification as complying with a standard set by non-governmental organizations, applied for by institutions, programs, and facilities on a voluntary basis] | http://id.nlm.nih.gov/mesh/D000068
Adequacy | Suitability of resources, processes, or systems to achieve intended objectives. In ISO 15189, adequacy is judged against documented requirements. | The adequacy of the PCR thermocycler and reagents was verified to ensure accurate amplification across all samples. | In bioinformatics, adequacy refers to whether computational infrastructure (e.g., storage, CPU/GPU resources), methods, pipelines, and reference datasets are sufficient to perform analyses reliably and accurately. For example, assessing whether available compute resources can handle whole-genome sequencing data or whether reference databases are comprehensive enough to detect updated resistance mutations. | The reference genome and variant database were reviewed for adequacy to ensure they included updated known resistance mutations in pathogen datasets, enabling accurate detection and reporting.
Authorisation | Official permission granted by a competent person or authority to perform a specific activity, release results, or operate equipment in accordance with established procedures. | The clinical scientist provided authorisation to release patient RNA-seq results after verifying QC metrics and reviewing the final report.
Calibration | Operation establishing the relationship between instrument indications and reference standards (ISO VIM). Calibration ensures that measurements are accurate, traceable, and reliable. | Calibration of the qPCR instrument was performed using standard reference materials to ensure accurate viral load quantification. | In bioinformatics, calibration focuses on benchmarking algorithms or tools against reference datasets rather than hardware. Examples include calibrating variant quality scores, adjusting machine-learning thresholds, or normalizing QC metrics using gold-standard datasets to ensure accurate and reliable results. | Calibration of the variant quality score model was performed using Genome in a Bottle reference calls to ensure accurate variant classification. | mesh:D002138 [Determination, by measurement or comparison with a standard, of the correct value of each scale reading on a meter or other measuring instrument; or determination of the settings of a control device that correspond to particular values of voltage, current, frequency or other output] ¦ NCIT:C69187 [The act of checking or adjusting (by comparison with a standard) the accuracy of a measuring instrument] | http://id.nlm.nih.gov/mesh/D002138 ¦ http://purl.obolibrary.org/obo/NCIT_C69187
Change Control | A formal process within the Quality Management System (QMS) for managing modifications to documents, procedures, equipment, or systems. It ensures that any changes are reviewed, approved, implemented, and documented in a controlled manner to maintain quality and compliance. | Missing | The formal management of changes to computational workflows, analysis pipelines, software tools, reference genomes, or databases. Change control ensures that updates are systematically reviewed, validated, approved, and documented before implementation, to prevent unintended impact on results and maintain reproducibility. | Change control was required when updating the reference genome for Mycobacterium tuberculosis to a new version, ensuring that the updated reference, pipeline re-validation, and impact on drug-resistance mutation reporting were fully documented before implementation.
Change Request (Document Control) | A formal request submitted to modify a controlled document (e.g., SOP, validation protocol). Change requests must include justification, details of the proposed modification, and undergo review and approval before implementation. All approved change requests must be recorded in the Quality Management System (QMS). | A change request was submitted to update the DNA extraction SOP to use a new extraction kit | See traditional definition | A change request was submitted to update the server IP address specified in the SOP following a change in the server configuration.
Checklists (under Audits) | Structured tools used to verify compliance with requirements, procedures, or standards during audits. Checklists help ensure that all critical elements are reviewed consistently. | A checklist was used during the internal audit to confirm that sample labeling, reagent storage, and instrument maintenance complied with SOPs. | See traditional definition | During pipeline audits, a checklist ensured that software versions, reference databases, QC metrics, and documentation were reviewed for compliance with quality standards. | mesh:D057189 [Aid for consistent recording of data such as tasks completed and observations noted.] | http://id.nlm.nih.gov/mesh/D057189
Clinical Accuracy | Degree to which test results correctly reflect the clinical status of the patient (ISO 15197, CLSI). Clinical accuracy goes beyond analytical correctness to measure how well results correspond to actual patient condition. | Clinical accuracy of the HIV viral load assay was confirmed by comparing measured values with patient outcomes and treatment response | In bioinformatics, clinical accuracy applies to diagnostic pipelines, ensuring that results reflect true patient status—for example, detecting a pathogen in sequencing data corresponds to an actual infection. It evaluates the clinical relevance of bioinformatics analyses beyond analytical performance metrics. | Clinical accuracy of the TB genomics pipeline was confirmed by comparing detected Mycobacterium tuberculosis sequences and resistance markers with patient microbiology results, ensuring reported variants reflected true infection and drug resistance status.
Clinical Decision Limits | Predefined values of a biomarker or measurement used for clinical decision-making (ISO 15189). | A clinical decision limit of 20 copies/mL, corresponding to the assay’s lower limit of quantitation, was applied for reporting HIV viral load; results below this threshold were reported as “detected but not quantifiable.” | In bioinformatics, clinical decision limits are predefined thresholds applied to computational results—such as variant allele frequencies, read depths, or genome assembly metrics—to guide reporting and clinical interpretation. Variants or signals below these limits are not reported because they cannot be reliably distinguished from errors or background noise. | A clinical decision limit was applied to variant allele frequencies, where only variants present above 20% were reported, as lower-frequency variants could not be reliably distinguished from sequencing or pipeline errors.
Commutability | Property of a reference material to behave like patient samples across different measurement procedures, ensuring that results are comparable and meaningful. | Reference plasma samples showed commutability by producing results consistent with patient specimens across multiple viral load assays. | In bioinformatics, commutability refers to whether reference datasets behave comparably to real patient samples across analysis pipelines. For example, synthetic spike-ins were assessed for commutability to ensure they accurately mimicked biological variability. | The synthetic reference dataset lacked commutability and did not represent patient sample variability, highlighting limitations for benchmarking variant detection pipelines.
Comparability | Ability to compare results across different laboratories, methods, or time points, within defined limits of agreement. Ensures that measurements are consistent and interpretable in different contexts. | Comparability of viral load results was confirmed by re-analyzing the same patient samples in three independent laboratories using the same SOPs. | In bioinformatics, comparability refers to ensuring that results from different pipelines, datasets, or centres can be aligned and interpreted consistently. This is achieved through standardized data formats, shared reference datasets, harmonized QC metrics, and agreed cut-offs or settings for computational tools. | Comparability of variant calling results across three sequencing centres was ensured by using shared reference datasets, harmonized QC metrics, and agreed tool settings and cut-offs
Competence | Demonstrated ability to apply knowledge, skills, and experience to achieve intended results (ISO/IEC 17025:2017; ISO 15189:2022). Includes performing assays, maintaining instruments, following SOPs and safety/QC protocols, and executing analyses with awareness of limitations. Competence must be documented and maintained, e.g., through training records, competency assessments, DOPs, and competency profiles (see competence record). | Laboratory staff demonstrated competence by planning and performing PCR assays, conducting QC checks, operating instruments correctly, and interpreting results while adhering to SOPs and acknowledging assay limitations. | In bioinformatics, competence is the ability to design, execute, and interpret analyses—such as quality control, pipeline selection, statistical testing, and data interpretation—while understanding the limitations of the tools, datasets, and methods used. | Competence was demonstrated by designing a new pathogen genomics pipeline to meet clinical requirements, selecting appropriate analysis tools, performing QC checks, and interpreting results while considering the limitations of datasets and algorithms. | OCCO:00000121 [A disposition borne by an organism in virtue of training such that, if realized, is realized in the successful performance of a skilled task for which that training was pursued] | http://purl.obolibrary.org/obo/OCCO_00000121
Competence record | Demonstrated ability to apply knowledge, skills, and experience to achieve intended results (ISO/IEC 17025:2017; ISO 15189:2022). Includes performing assays, maintaining instruments, following SOPs and safety/QC protocols, and executing analyses with awareness of limitations. Competence must be documented and maintained, e.g., through training records, competency assessments, direct observation of procedures (DOPs), and competency profiles. | Laboratory staff competence was demonstrated through DOPs and competency profiles, showing accurate execution of laboratory procedures—including extraction, library preparation, and assay setup—proper instrument maintenance, adherence to SOPs, and compliance with safety and QC protocols. | In bioinformatics, competence is the demonstrated ability to apply knowledge, skills, and experience to achieve intended results in computational analyses. This includes designing and executing analyses, developing and running pipelines, managing software, reference datasets, and computing infrastructure (hardware and storage), performing quality control checks, and interpreting results with awareness of the limitations of tools, methods, and datasets. Competence must be documented and maintained, e.g., through training records, competency assessments, direct observation of procedures (DOPs), and competency profiles. | Competence in bioinformatics was demonstrated by the ability to perform routine computational tasks—including pipeline execution, data preprocessing, QC checks, and result interpretation—while managing software, reference datasets, and computing infrastructure, as assessed through documented observation of procedures (DOPs), competency profiles, and formal training records.
Consensus Data | Results agreed upon by multiple experts, methods, or laboratories, often serving as a benchmark. In ISO, consensus data underpin reference values or validation and must be determined systematically. | Consensus data from three independent laboratories were used to establish the reference range for viral load measurements. | In bioinformatics, consensus data can be a gold-standard variant set derived from multiple pipelines or jointly curated annotations, used to validate new methods or resolve ambiguous results. For example, the Genome in a Bottle (GIAB) variant set serves as consensus data for benchmarking human variant-calling pipelines. | The TB variant-calling pipeline was benchmarked against a consensus dataset compiled from multiple sequencing centers to ensure accuracy and reproducibility of detected resistance mutations.
Contamination | Unintended presence of material, organism, or signal that compromises results. | Contamination was detected when non-template controls showed higher-than-expected read counts, indicating cross-sample contamination during nucleic acid extraction. | In bioinformatics, contamination refers to unwanted reads, cross-sample index misassignments, or sequences from off-target organisms that can affect analysis. | Contamination was detected when off-target pathogen reads, such as influenza sequences, appeared in SARS-CoV-2 samples, indicating cross-sample contamination. | PATO:0015031 [The presence of a constituent, impurity, or some other undesirable element that spoils, corrupts, infects, makes unfit, or makes inferior a material, physical body, natural environment, place of human occupancy, or other material entity] | http://purl.obolibrary.org/obo/PATO_0015031
Controls (Positive / Negative) | Specimens with known characteristics used to verify that a test, assay, or procedure is performing correctly. Positive controls contain the target analyte and should yield a positive result, while negative controls lack the analyte and should yield a negative result. Controls help identify assay failures, contamination, or procedural errors. | Positive and negative controls were included in each DNA extraction run to verify extraction efficiency and ensure the absence of contamination. | In bioinformatics, controls are datasets used to verify pipeline performance and detect errors. Positive controls include datasets with expected variants or signals, which can be real experimental data or in silico–generated datasets simulating expected features. Negative controls, such as empty libraries, no-template runs, or in silico blanks, are used to detect contamination or spurious results. | Positive control datasets containing known SNVs were used to confirm variant calling accuracy, while no-template runs served as negative controls to detect contamination. | There are ontologies for positive and negative controls - but not for the term controls
Correction Factors | Factors applied to correct for known systematic biases or effects in measurements, ensuring that reported values more accurately reflect the true quantity or state. | Correction factors were applied to qPCR results to adjust for variations in viral load measurements due to instrument drift and sample dilution. | See plain definition | Correction factors were applied to sequencing data to normalize read counts (e.g., TPM/RPKM), adjust for batch effects, and correct error rates in variant calling.
Corrective Action | Action taken to eliminate the cause of a detected non-conformance or other undesirable situation. The aim is to prevent recurrence rather than simply address the immediate issue. | Corrective actions address the root cause of failed QC results, contamination events, or SOP deviations — for example, retraining staff or revising procedures after a control failure. | In bioinformatics, corrective actions may be triggered by pipeline failures, untraceable results, or deviations from documented workflows. They go beyond fixing the symptom — addressing the underlying cause through procedural updates, training, or improved documentation. | A corrective action was implemented by updating the SOP to require version-controlled reference genomes to ensure traceability and reproducibility.
Cross Audit | A peer-to-peer audit carried out between collaborating organisations or departments to ensure consistent application of quality standards and identify areas for improvement. | A cross audit between two hospital laboratories compared sample handling and reporting procedures to align best practices. | See traditional definition | A cross audit between two genomic centres reviewed variant-calling pipelines to ensure consistent validation and traceability across sites.
Development | A structured process of designing, creating, and refining new methods, assays, systems, or workflows. Under ISO 15189, development must follow controlled and documented stages — including design input, design output, verification, validation, and implementation — to ensure the final product or method meets defined performance specifications and intended use. | Development of a new RNA extraction protocol was carried out under design control, with verification against existing methods and validation using clinical samples to confirm performance suitability. | A structured process of designing, creating, and refining new methods, assays, systems, or workflows. Under ISO 15189, development must follow controlled and documented stages — including design input, design output, verification, validation, and implementation — to ensure the final product or method meets defined performance specifications and intended use. In bioinformatics, development often follows agile software development principles. | Development of the HIV bioinformatics pipeline was performed under design control, with verification using synthetic datasets and validation against reference genomes to ensure accurate detection of known resistance variants and reliable prediction of their clinical impact.
Deviations | A documented departure from an approved laboratory procedure, specification, or expected result. Deviations may arise from skipped steps, use of incorrect reagents, instrument malfunction, or environmental control failures. All deviations must be recorded in the QMS, investigated for root cause, and addressed through corrective or preventive actions (CAPA). | A deviation was recorded when an extraction was performed using a reagent lot that had not been approved under the current SOP. | A documented departure from an approved computational procedure, specification, or expected outcome. Examples include use of unvalidated software versions, incorrect reference genomes, or failure to follow the defined pipeline workflow. Deviations must be recorded, investigated, and resolved through CAPA to maintain reproducibility, traceability, and compliance. | A deviation was recorded when the analyst used a reference genome build that was not specified in the validated pipeline SOP.
Discrepancy | Documented difference between observed and expected results, processes, or records. | The measured DNA concentration of a sample differed from the value recorded during initial quantification, triggering a review of sample handling and pipetting | A documented difference between expected and observed results, workflows, or data. This includes mismatches between SOP instructions and actual pipeline parameters, between expected QC values and observed results, or between repeated pipeline runs. | Pipeline parameters did not match SOP instructions, QC metrics fell outside expected thresholds, or reruns of the same analysis produced differing variant counts.
Document Control | A formal process to ensure documents are reviewed, approved, distributed, updated, and archived within a quality management system (QMS). | Document control procedures ensured obsolete DNA extraction SOPs were archived and only current versions were used. | A formal process to manage and track pipeline documentation, including SOPs, scripts, and configuration files, ensuring only approved and current versions are used. | Document control procedures ensured the SOP was updated to include the latest validated version of a specific pipeline.
External Audit | An audit conducted by an independent organisation, such as an accreditation body or regulatory authority, to verify compliance with standards (e.g., ISO 15189, ISO 17025). | An external audit by UKAS assessors confirmed that all testing processes met ISO 15189 requirements. | See traditional definition | An external ISO audit included review of computational processes, confirming that data handling, validation, and documentation met accreditation standards.
External Quality Assessment (EQA) | Formal evaluation of a laboratory’s or analytical service’s performance through participation in external comparison programmes such as UK NEQAS or equivalent schemes. EQA verifies that internal quality control procedures, methods, and analytical outputs remain accurate, reliable, and consistent with external benchmarks. | The laboratory participated in a UK NEQAS Molecular EQA scheme for Mycobacteria molecular detection and resistance testing, processing simulated sputum samples to assess detection accuracy and rifampicin resistance prediction. Participation verified the lab’s performance against national benchmarks. | Evaluation of a bioinformatics pipeline or analysis by comparing outputs with external datasets, peer labs, benchmarking challenges, or gold-standard datasets. | The bioinformatics team participated in an inter-laboratory EQA exercise by re-analyzing datasets from the UK NEQAS Molecular SARS-CoV-2 variant typing scheme and comparing variant calls against gold-standard references. This confirmed pipeline accuracy for variant identification.
Finding | Result of an evaluation against requirements, identifying conformity, non-conformity, or opportunities for improvement. | An audit finding noted that sample storage temperatures were not consistently recorded. |  | An audit finding noted that pipeline versions were not recorded in the report. | NCIT:C61036 [The result of an examination or inquiry] | http://purl.obolibrary.org/obo/NCIT_C61036
Functional Specification | A document detailing what a system, process, or component must achieve, without prescribing how it should be done. | The functional specification for a DNA extraction process required consistent yield and purity across sample types. | A document specifying the expected inputs, outputs, performance, and features of a pipeline or software tool, without dictating implementation. | The functional specification for the pipeline required it to accept FASTQ input and produce annotated VCF output within defined runtime parameters.
Improved Performance | Demonstrated enhancement in the efficiency, effectiveness, or reliability of a process following a change. Improvement must be measured against predefined indicators. | Improved performance was demonstrated by shorter turnaround times and reduced sample handling errors after workflow optimization. | Demonstrated enhancement in the efficiency, accuracy, or reliability of a bioinformatics pipeline or analysis following a change. Improvement must be measured against predefined metrics. | Improved performance was demonstrated by reducing false positive variant calls by 20% and adding new functionality for structural variant detection after pipeline optimization.
Incident | An occurrence that led or could lead to non-conformance, error, or deviation from expected results. | An incident occurred when a sample was mislabeled, resulting in an incorrect test assignment and an out-of-TAT result. | See traditional definition | An incident occurred when a misconfigured pipeline parameter caused incorrect variant calls across multiple samples.
Inconsistencies | Lack of uniformity or agreement in data, documents, or results. ISO expects identification and resolution of inconsistencies. | Inconsistencies were found between duplicate PCR reactions, prompting a review of sample handling procedures. | See traditional definition | Inconsistencies were identified between pipeline documentation and the parameters actually applied.
Installation Qualification (IQ) | Documented verification that equipment is installed correctly and meets the manufacturer’s specifications. Part of the validation process (IQ/OQ/PQ). | Installation qualification confirmed that a new PCR machine was set up according to the manufacturer’s instructions. | Documented verification that software, pipelines, and supporting compute infrastructure are installed correctly and meet expected specifications. Part of the validation process (IQ/OQ/PQ). | Installation qualification confirmed that the pipeline and HPC environment, including all dependencies and compute nodes, were correctly configured.
Interferences | Influence from a substance, signal, or condition that alters the true measurement. | Interferences from primer-dimers or nonspecific amplification affected qPCR quantification accuracy. | Influence from contaminating sequences, technical artefacts, cross-talk, or algorithmic bias that distorts analysis results. | Interferences from index hopping introduced artefactual reads into the dataset. | CAO:000046 [A systematic error in the measure of a signal caused by the presence of concomitants in a sample] | http://champ-project.org/images/ontology/cao.owl#CAO_000046
Internal Audit | A systematic, independent evaluation conducted within an organisation to determine whether activities and related results comply with planned arrangements and meet quality management system (QMS) requirements. | An internal audit reviewed sample tracking and equipment calibration logs to confirm compliance with the laboratory’s quality procedures. | See traditional definition | An internal audit reviewed pipeline traceability, code version control, and validation records to ensure compliance with documented SOPs.
Internal Quality Control | Procedures performed within the laboratory’s routine operations to monitor ongoing performance and detect errors in real time. Often involves control samples run alongside test samples. | Daily instrument controls, reagent blanks, and positive/negative controls were included in each qPCR run. | Procedures within routine data analysis to monitor pipeline performance and detect errors or anomalies in real time. | QC metrics and control datasets were analyzed alongside sample data to identify unexpected coverage drops or variant calling errors.
Language | System of terms, symbols, or structured vocabulary used for communication. In ISO contexts, clarity and consistency of language is critical for compliance. | The laboratory ensured that terms like “validation” and “verification” were used consistently across SOPs and reports, maintaining clarity and traceability. | In bioinformatics, language has dual meaning: human language (terminology in SOPs, ontologies) and programming language (R, Python). ISO focus is on controlled terminology, so “language” here means ensuring words are used consistently across SOPs and data dictionaries.
Management System | A set of interrelated or interacting elements used to establish policies, objectives, and processes to achieve them. In laboratories, this typically includes the Quality Management System (QMS) and can extend to cover data analysis, documentation, software versioning, traceability, and validation/verification of pipelines. | The management system was updated to include procedures for sample tracking and instrument calibration. | See traditional definition | The management system was updated to include documented procedures for software version control and data retention in bioinformatics workflows
Method | A systematic approach or technique used to achieve a specific objective. | The qPCR amplification method was chosen to quantify SARS-CoV-2 viral RNA accurately. | NA | The read alignment method using bwa mem was applied to map sequencing reads to the reference genome. | NCIT:C71460[A means, manner of procedure, or systematic course of actions that have to be performed in order to accomplish a particular goal] | http://purl.obolibrary.org/obo/NCIT_C71460
Metrological Traceability | In wet labs, metrological traceability is the gold standard for linking measurements (e.g., concentration) back to SI units through reference materials and calibrated instruments. | Metrological traceability of viral load estimates was established using reference datasets linked to WHO international standards. | In bioinformatics, metrological traceability refers to linking computed or digital results (e.g., expression levels, variant frequencies) back to reference standards, validated datasets, or algorithms, with documented uncertainty, thereby anchoring digital outputs to real-world measurement systems. | Variant allele frequencies in pathogen genomics were benchmarked against Genome in a Bottle (GIAB) reference datasets to establish metrological traceability of computational measurements.
Near-miss | An occurrence that could have caused harm or non-conformance but did not. | A near-miss occurred when two samples were swapped, but the error was caught during the procedure. | See traditional definition | A near-miss occurred when an outdated HIV resistance database was nearly used, which would have missed a clinically relevant variant in a sample, but the issue was detected before reporting results.
Noise | Unwanted variation or disturbance that obscures the measurement of the true signal. Noise can be random or systematic and may arise from instrument background signals, contamination, sample degradation, or environmental fluctuations. | Background fluorescence from the plate reader contributed noise that was accounted for during viral load quantification. | Random sequencing errors or background reads not associated with the true signal (e.g., low-quality base calls, index hopping). Noise can also appear as spurious variation in high-dimensional datasets (e.g., dropout in single-cell RNA-seq). | The signal-to-noise ratio was evaluated in SARS-CoV-2 sequencing data to filter out low-quality reads and minimize spurious variant calls | NCIT:C45783 [Any part of a signal that is not the true or original signal but is introduced by the communication mechanism] | http://purl.obolibrary.org/obo/NCIT_C45783
Non-conformance | Non-fulfilment of a requirement. May arise from deviations, errors, or failures to meet specifications. | A non-conformance was recorded when the laboratory technician deviated from the SOP during RNA extraction, potentially affecting downstream sequencing results. | See plain definition | A non-conformance was identified when a non-validated reference database was used for pathogen genome annotation.
Patient Safety | Reduction of risk of unnecessary harm to patients to an acceptable minimum (ISO 15189). | Patient safety is maintained by following validated protocols, including proper labeling and handling, and using appropriate controls to prevent incorrect results or sample mix-ups that could lead to misdiagnosis. | See plain definition | In clinical bioinformatics, patient safety involves ensuring that analyses and reports do not cause harm through misclassification, delays, or unclear communication, such as false negatives in pathogen detection.
Performance Qualification (PQ) | Evidence that systems/processes perform as intended under routine (real-world) conditions. | Performance qualification demonstrated that the sequencing pipeline maintained processing of ≥100 samples per day with stable accuracy and no missed service-level agreements (SLAs) over 30 consecutive days. | See plain definition | For computational workflows, PQ demonstrates that a pipeline and infrastructure deliver consistent accuracy and throughput under production load with monitored stability. While the term “PQ” is rarely used outside regulated settings, the principle ensures reliable bioinformatics performance in routine use.
Performance Reporting | Structured communication of performance against defined objectives or indicators, often using agreed metrics to track quality, efficiency, and reliability. | Quarterly performance reporting summarized contamination rates, turnaround times, and repeat extraction frequencies against established targets. | Performance reporting in bioinformatics may take the form of periodic dashboards or reports summarizing key performance indicators such as mapping rate, turnaround time, re-analysis rate, pipeline failure rates, and variant calling accuracy, with trends highlighted and corrective actions documented.
Performance Review | Formal evaluation of performance against predefined objectives, standards, or indicators, used to identify areas for improvement and ensure ongoing compliance with quality requirements | The semi-annual performance review identified missed turnaround time (TAT) targets and triggered a capacity upgrade. | See plain definition | A performance review of the pathogen genomics pipeline identified recurring pipeline failures and suboptimal mapping rates, leading to revalidation and infrastructure improvements as part of CAPA.
Performance Verification | Confirmation that a system, method, or pipeline performs according to predefined requirements or specifications. Verification demonstrates that stated performance characteristics are achieved, but is narrower in scope than full validation (which assesses overall suitability for intended use). | Performance verification confirmed that the PCR assay consistently detected target pathogens at the predefined limit of detection. | See plain definition | Performance verification demonstrated that the pathogen genomics pipeline achieved the stated runtime and reproducibility, with sensitivity ≥95% and specificity ≥99% based on the Genome in a bottle (GIAB) test set.
Preparation | Process of making something ready for use, often referring to specimens, reagents, or documents. Steps are controlled and documented to ensure reproducibility and traceability. | Preparation involved creating reagent aliquots, labeling samples, and setting up specimens according to SOPs. | Process of making data ready for analysis, including file formatting, quality control filtering, adapter trimming, and metadata curation. Steps must be controlled and documented to ensure reproducibility. | Preparation of the sequencing dataset included adapter trimming, QC filtering, file formatting, and metadata standardization before downstream analysis.
Preventive Actions | Actions taken to eliminate the cause of a potential non-conformance, before it occurs. Preventive actions are proactive measures designed to reduce risk and improve system robustness. | As a preventive action, the laboratory introduced dual sample labeling and independent verification steps to avoid potential mix-ups during DNA extraction. | See plain definition | As a preventive action, we added pre-release smoke tests and immutable reference files to avoid silent regressions in the pathogen genomics pipeline. Other examples include enforcing version pinning, implementing storage redundancy, and using continuous integration (CI) checks to prevent configuration drift or silent failures.
Procedure | A defined series of steps to implement a method, usually documented in a Standard Operating Procedure (SOP). | The SOP outlined the RNA extraction and cDNA synthesis procedure for all clinical samples. | Step-by-step instructions for performing a specific aspect of a bioinformatics workflow, with defined inputs, parameters, expected outputs, and traceable records. | The QC procedure details how raw sequencing reads are checked for quality metrics, trimmed, adapters removed, deduplicated, and filtered before downstream analysis, with all steps and parameters documented to ensure reproducibility and traceability. | SIO:000999 [A procedure is a process that attempts to achieve one or more objectives by following an established set of actions] ¦ NCIT:C79751 [Any stepwise set of actions] | http://semanticscience.org/resource/SIO_000999 ¦ http://purl.obolibrary.org/obo/NCIT_C79751
Process | Series of interrelated or interacting activities that transform inputs into outputs. | The DNA extraction and library preparation process converts tissue or blood samples into purified, sequencer-ready DNA libraries. | NA | The variant analysis process converts raw FASTQ reads into annotated VCF files through sequential steps including mapping, deduplication, variant calling, and variant interpretation. | NCIT:C29862 [A sequence of events which produces some outcome] | http://purl.obolibrary.org/obo/NCIT_C29862
Quality | Degree to which a set of inherent characteristics of an object fulfills requirements. It reflects fitness for intended use under a Quality Management System (QMS). | The quality of the RNA extraction was assessed by measuring yield, purity, and integrity against predefined acceptance criteria. | Degree to which a computational analysis or dataset meets defined criteria for accuracy, reliability, and reproducibility under a QMS. | The quality of the sequencing analysis was evaluated by comparing read depth, base quality scores, duplication rates, and alignment metrics to predefined acceptance thresholds. | NCIT:C25283 [An essential and distinguishing attribute of something or someone; a degree or grade of excellence or worth] | http://purl.obolibrary.org/obo/NCIT_C25283
Quality control | Operational techniques and activities used to fulfil quality requirements, ensuring results are consistent, valid, and reliable. Includes checks on reagents, instruments, calibration, and control samples, performed according to documented procedures and predefined thresholds. | QC of RNA extraction included monitoring reagent performance, running positive and negative controls, and checking instrument calibration. | Operational checks applied to sequencing or analysis data to ensure accuracy and reliability, anchored to predefined, measurable criteria rather than informal filtering. May include adapter trimming, base-quality filtering, duplicate removal, and assessment of mapping rates. | QC of RNA-seq reads included trimming adapters, filtering low-quality bases, removing duplicates, and evaluating mapping rates to confirm data suitability for downstream analysis. | mesh:D011786 [A system for verifying and maintaining a desired level of quality in a product or process by careful planning, use of proper equipment, continued inspection, and corrective action as required] ¦ NCIT:C15311 [Set of measurements and inspections taken to verify that performance of equipment and procedures remains within specified limits] | http://id.nlm.nih.gov/mesh/D011786 ¦ http://purl.obolibrary.org/obo/NCIT_C15311
Quality Indicators | Quantitative measures used to monitor and evaluate how well a process meets predefined quality objectives. Indicators should be measurable, predefined, and regularly reviewed to support continuous improvement. | Quality indicators in the laboratory included turnaround times, error rates, and frequency of repeat testing. | See plain definition | In bioinformatics, quality indicators may include read mapping rates, error rates in variant calls, reproducibility scores, and user satisfaction with reports. These indicators should be predefined, measurable, and regularly reviewed to ensure pipeline performance and reliability.
Quality Management System (QMS) | A structured set of policies, processes, procedures, and responsibilities used by an organization to ensure quality objectives are achieved, risks are managed, and compliance with applicable standards is maintained (ISO 9000/ISO 15189). | The laboratory implemented a QMS encompassing SOPs, document control, internal audits, equipment calibration, and staff training to ensure reliable and reproducible molecular test results.
Reference | A source that provides evidence or context for a decision, requirement, or comparison. In ISO, often refers to authoritative standards or datasets. | Wet-lab references include international standards, reference ranges, or cited procedures used to interpret test results. | In bioinformatics, reference materials include reference genomes, databases, or curated datasets used as benchmarks. | Variant calls were compared against HIV references and HIV variant databases to detect known resistance mutations and validate accuracy.
Reference Material | Material that is sufficiently homogeneous and stable with respect to specified properties, established to be fit for its intended use in measurement. Certified Reference Materials (CRMs) have documented property values with stated uncertainty. Controls run within an experiment are not reference materials; only external standards with traceable properties qualify. | A certified reference DNA sample was used to verify assay accuracy across multiple qPCR runs. | Material or datasets used to validate and verify bioinformatics pipelines, providing a known baseline for comparison. This includes gold-standard datasets (e.g., NA12878 genome from GIAB), well-characterized cell lines, or spike-in controls. | The sequencing pipeline was verified using a Genome in a Bottle GIAB reference genome to confirm variant calling accuracy. | CHMO:0002913 [A substance or mixture of substances, the composition of which is known within specified limits, and one or more of the properties of which is sufficiently well established to be used for the calibration of an apparatus, the assessment of a measuring method or for assigning values to materials] | http://purl.obolibrary.org/obo/CHMO_0002913
Reporting | Formal communication of laboratory test results in a controlled and traceable format, ensuring clarity, accuracy, and compliance with the Quality Management System (QMS). | Reporting included generating a PDF report for clinicians summarizing HIV viral load, QC metrics, and validated assay results for resistance testing, enabling informed treatment decisions. | Controlled communication of analysis results in a reproducible and traceable format. Includes QC summaries, variant calls, predicted phenotypes, and structured outputs for stakeholders. May involve submission of variants to international databases. Ensures results are interpretable, accurate, and auditable within the QMS. | Reporting of HIV sequencing data included resistance variants and predicted susceptibility to antiretroviral therapies, delivered as structured outputs (e.g., JSON for databases, PDF for clinicians), and included submission of variants to international databases such as HIVdb (Stanford HIV Drug Resistance Database).
Requirement for Reports | ISO requires that reports include minimum information to ensure traceability and interpretability, such as sample/patient identification, methods, results, interpretations, and signatures. | The requirement for reports mandated inclusion of sample identification, assay method, QC control results, measured viral load, and signature of the clinical scientist, ensuring results were traceable, verified, and suitable for clinical decision-making. | ISO requires that computational analysis reports include minimum information such as software version, reference genome, pipeline parameters, QC metrics, and results interpretations to ensure reproducibility and compliance with the QMS. | The requirement for reports mandated inclusion of software version and reference genome used, along with QC metrics, in the sequencing report.
Result Validity | Extent to which a result can be considered sound, accurate, and supported by evidence. In ISO, validity is linked to adherence to validated methods, QC, and traceability. | Result validity was ensured by verifying assay outputs against a certified reference material and confirming QC metrics were within acceptable ranges. | See traditional definition | Result validity is demonstrated when all computational analyses produce consistent, accurate, and reproducible outputs, using validated pipelines, appropriate reference data, and quality-controlled input, so that the results can be confidently used for clinical or public health decisions.
Review | Systematic examination of activities, results, or documents against requirements. In ISO, a review means a documented, traceable sign-off that ensures accuracy, completeness, and compliance. | All sequencing reports underwent independent review, with a second clinical scientist performing a counter-sign check to confirm the analysis before release. | In bioinformatics, review applies to SOPs, pipeline code, and analysis reports, including code reviews and validation of outputs. Traceability is maintained via review logs, approvals, and audit trails to demonstrate compliance with validated standards and reproducibility. | The updated influenza pipeline underwent code review by a second bioinformatician before deploying the updated pipeline.
Risk Analysis | Systematic use of information to identify sources of risk and estimate their likelihood and impact (ISO 14971, ISO 31000). | Risk analysis identified hazards in sample handling, including potential contamination during extraction, mislabeling of samples, degradation during transport, and improper storage conditions, allowing the laboratory to implement mitigation measures. | Risk analysis assesses risks in computational workflows, including pipeline failures, data loss, unauthorized access, misinterpretation of results, and algorithmic biases. It guides the implementation of risk controls and informs decision-making. | Risk analysis identified the possibility of pipeline failure or data loss due to insufficient storage redundancy, leading to backup systems and version-controlled pipelines. | mesh:D018570 [The qualitative or quantitative estimation of the likelihood of adverse effects that may result from exposure to specified health hazards or from the absence of beneficial influences] | https://meshb.nlm.nih.gov/record/ui?ui=D018570
Risk Management | Risk management involves identifying, evaluating, and mitigating hazards in laboratory processes, including sample handling errors, contamination, instrument failure, and assay deviations. Controls may include standardized SOPs, staff training, QC procedures, and maintenance schedules. | Risk management in the laboratory included second-person checks to reduce the risk of sample mix-ups during lab procedures. | Risk management covers pipeline errors, software bugs, incorrect reference databases, misannotation, or reproducibility failures, and extends to data security and patient confidentiality when handling clinical data. Controls include version-controlled pipelines, automated QC checks, backups, and access restrictions. | Risk management for the updated influenza pipeline included version-controlled code review and automated unit testing/continuous integration (CI) testing to reduce the risk of negative impact on results, such as producing erroneous or misleading outputs once deployed. | mesh:D012308 [The process of minimizing risk to an organization by developing systems to identify and analyze potential hazards to prevent accidents, injuries, and other adverse occurrences, and by attempting to handle events and incidents which do occur in such a manner that their effect and cost are minimized.]
Sample/specimen | One or more items taken from a population or individual and intended to provide information about that population or individual (ISO 15189 / ISO 17025). In wet labs, this is usually the patient specimen or aliquot under test, such as whole blood, tissue biopsy, swab, urine, stool, or other biological material. | Each sample was assigned a unique identifier before sequencing to ensure traceability between the physical specimen and associated laboratory records. | In bioinformatics, sample can also refer to a dataset derived from a single biological specimen, such as FASTQ files, BAM files, or cell counts. In multi-omics studies, one specimen may generate multiple data types (DNA, RNA, protein), and careful tracking is needed to maintain traceability. | For HIV resistance analysis, each sample dataset—including FASTQ and BAM files derived from a patient’s blood specimen—was tracked with a unique identifier, ensuring traceability and reproducibility across the pipeline, even when multiple data types (DNA, RNA) were generated from the same specimen. | NCIT:C19157 [A part of a thing, or of several things, taken to demonstrate or to determine the character of the whole, e.g. a substance, or portion of material obtained for use in testing, examination, or study; particularly, a preparation of tissue or bodily fluid taken for examination or diagnosis] ¦ SIO:001050 [A sample is a limited quantity of something (e.g. an individual or set of individuals from a population, or a portion of a substance) to be used for testing, analysis, inspection, investigation, demonstration, or trial use] | http://purl.obolibrary.org/obo/NCIT_C19157 ¦ http://semanticscience.org/resource/SIO_001050
Sample Quality | Degree to which a sample is suitable for intended testing, including integrity, stability, and absence of contamination. | Sample quality was assessed by measuring RNA integrity numbers (RIN) and total concentration to confirm suitability for sequencing. | Sample quality reflects whether sequencing data derived from a specimen is reliable. Metrics include read length, duplication rate, adapter contamination, and overall sequence integrity, guiding decisions for downstream analysis and reporting. | Sample quality was evaluated by examining read length, duplication rate, adapter contamination, and overall sequence integrity to ensure reliable input for variant calling. | SIO:001411 [sample quality is the quality of an object that is drawn from a larger population] | http://semanticscience.org/resource/SIO_001411
Scope | Extent and boundaries of activities, analyses, or services covered by a document, SOP, or accreditation. | The scope of this SOP covers DNA sequencing workflows using Illumina platforms only. | See traditional definition | The scope of this SOP covers analysis of Illumina DNA sequencing data, including quality control, alignment, and variant calling.
Selection | Choosing appropriate methods, materials, or procedures to meet requirements. | The laboratory documented the selection of an RNA extraction kit suitable for blood and tissue samples, with justification based on yield and purity requirements. | In bioinformatics, selection means choosing reference genomes, algorithms, or parameters appropriate for the intended use. Often tied to justification of why a method was chosen. | The selection of the GRCh38 reference genome was documented, including rationale for clinical suitability and compatibility with existing pipelines.
Selectivity | Ability of a method to distinguish the analyte of interest from other components. | The assay showed high selectivity by correctly distinguishing SARS-CoV-2 reads from background human RNA. | Ability of an analysis pipeline to detect the true signal in the presence of confounders, such as identifying a pathogen without false signals from host DNA. | The bioinformatics pipeline demonstrated high selectivity by correctly identifying viral reads while excluding host and environmental sequences.
Sensitivity | Ability of a method to correctly identify positives (true positive rate). | The qPCR assay achieved a sensitivity of 98% for detecting the target gene | Ability of a pipeline or algorithm to correctly identify true positives, expressed as the proportion of true variants or targets detected. | The sequencing pipeline achieved a sensitivity of 98% for SNP detection in the GIAB dataset. | OBCS:0000058 [A data item that measures the proportion of actual positives which are correctly identified as such] | http://purl.obolibrary.org/obo/OBCS_0000058
SI Units | International System of Units — globally agreed reference units for measurement (ISO/IEC 80000). | RNA concentration was measured as 50 ng/µL | SI units are used when linking digital results back to physical measurements, e.g., read depth reported as coverage per base pair, genome size in base pairs (bp), RNA yield in ng, or runtimes in seconds. Using SI units ensures interoperability and standardization across laboratories and pipelines. | Sequencing read depth was reported as coverage per base pair. | mesh:D007394[A system of physical units in which the fundamental quantities are length, time, mass, electric current, temperature, luminous intensity, and amount of substance, and the corresponding units are the meter, second, kilogram, ampere, kelvin, candela, and mole.] | http://id.nlm.nih.gov/mesh/D007394
Specificity | Ability of a method to correctly identify negative results (true negative rate). | The qPCR assay demonstrated 99% specificity, showing minimal cross-reactivity with non-target sequences. | Ability of a pipeline or algorithm to avoid false positives, correctly identifying true negatives. Applied in variant calling, classification, and pathogen screening. | The sequencing pipeline demonstrated 99% specificity for SNP detection in the GIAB dataset. | NCIT:C41395 [The probability that a test will produce a true negative result when used on non-effected subjects as compared to a reference or "gold standard". The specificity of a test can be determined by calculating: number of true negative results divided by the sum of true negative results plus number of false positive results] ¦ OBCS:0000059 [A data item that refers to the proportion of negatives in a binary classification test which are correctly identified] | http://purl.obolibrary.org/obo/NCIT_C41395 ¦ http://purl.obolibrary.org/obo/OBCS_0000059
Standard Operating Procedure (SOP) | Controlled document describing the approved way to perform a specific activity (ISO 9000). | The SOP for RNA extraction defines the protocol for isolating RNA from specific sample types, including reagents, incubation times, and yield assessment. The SOP was reviewed, approved, and version-controlled under document control. | In bioinformatics, SOPs govern data handling, QC, analysis, reporting, and archiving. Falls under document control: must be versioned, reviewed, approved, and archived. | The SOP for weekly system health checks defines procedures for monitoring software versions, disk usage, pipeline runtimes, and backup status. It is reviewed, approved, and version-controlled to ensure traceability and compliance with the QMS. | NCIT:C48443 [An established procedure to be followed in carrying out a given operation or in a given situation] | http://purl.obolibrary.org/obo/NCIT_C48443
Storage Space | Resources for preserving samples, records, or data safely and accessibly. | Freezer storage space and conditions were monitored regularly to ensure sufficient capacity and that samples were maintained under appropriate conditions. | In bioinformatics, storage space refers to digital storage for raw sequencing data, intermediate files, and results. ISO requires documented monitoring, capacity planning, and protection against data loss, such as regular backups and controlled access. | Storage space for sequencing data was monitored monthly to ensure compliance with retention and data integrity policies.
Suitability | Extent to which something is appropriate for its intended purpose. In ISO, suitability is often assessed during validation or verification, and in wet labs refers to whether materials, reagents, or reference standards are appropriate for their intended use. | The RNA extraction kit was evaluated for suitability with blood and tissue samples to ensure efficient recovery and integrity of RNA. | Suitability refers to whether a pipeline, dataset, or reference genome is appropriate for the intended computational analysis. | Suitability was ensured by choosing the appropriate reference genome based on the virus of interest and selecting a reference that met a closeness threshold relative to the sample sequence to maximize accuracy of variant calling.
Technical Knowledge | Evidence-based understanding of principles, methods, and standards required to perform a task competently. | Lab personnel performing next-generation sequencing must demonstrate technical knowledge of library preparation, sequencing assays, instrumentation, and associated quality control metrics to ensure reliable and reproducible results. | In bioinformatics, technical knowledge includes understanding algorithms, data formats, QC principles, software development practices, programming languages (e.g., Python, R), workflow management systems, and compute infrastructure. | During the development of a new analytical pipeline for SARS-CoV-2, the bioinformatician drew on their technical knowledge when selecting appropriate bioinformatics tools and reference genomes, implementing workflow management with Git, coding reproducible modules, and testing the pipeline using datasets representing both standard and edge cases to ensure reliability and reproducibility.
Traceability | Ability to trace the history, application, or location of an object or activity (ISO 9000). In metrology, often linked to the chain of calibrations to standards. | Traceability was maintained by logging each RNA sample’s extraction batch, operator, and reagent lot number, ensuring that any downstream result could be fully audited. | In bioinformatics, traceability refers to linking results back to raw data, pipeline versions, parameters, reference files, and analysts, enabling reproducibility and auditability. | Traceability of the variant call was maintained by linking the result to raw FASTQ files, pipeline version 3.2, and the reference genome build.
Training / re-training | Process of developing (training) or updating (re-training) knowledge, skills, and behaviours to ensure competence. In ISO contexts, training applies to personnel — ensuring staff are competent to perform assigned tasks according to current procedures. Re-training occurs following updates to SOPs, new technology implementation, or identification of nonconformances. | Staff were re-trained on the updated DNA extraction SOP following a change in reagent supplier to ensure consistent performance. | In ISO, training is for staff. In bioinformatics, “training” can also mean model training. Disambiguation is critical: training staff ensures competence in pipelines, while training models refers to machine learning. Re-training is triggered by SOP changes or nonconformance. | Analysts were re-trained following the introduction of a new variant calling pipeline to ensure consistent and correct usage across the team.
Test (Examination method / procedure / process) | A test (or examination, per ISO 15189) is the set of controlled operations that apply a defined method to a sample to generate a measurable result. In ISO terms, it represents the act of measurement or analysis performed under specified and validated conditions. | A PCR assay is a test that amplifies and detects target DNA sequences to confirm the presence of a pathogen in a patient sample. | In bioinformatics, a test refers to the analytical process that transforms raw sequencing or molecular data into interpretable results using defined computational methods. | Variant calling is a bioinformatics test that identifies genomic variants from aligned sequencing data using a validated algorithm and parameters.
Tolerance | Permissible range of variation in a measurement, value, or condition without invalidating the result (ISO VIM). Tolerance refers to acceptable deviations in experimental measurements or conditions that do not compromise the validity of results. | A tolerance of ±5% in pipetted reagent volumes was allowed to account for minor variations without affecting assay performance. | In bioinformatics, tolerance applies to predefined thresholds that results must stay within. Examples: acceptable mismatch rates in alignments, allowable run-time variance, or thresholds for QC metrics. It ensures outputs are still valid despite small variations. | A tolerance of ±5% in mapping rate was allowed between replicate sequencing runs, and QC thresholds defined acceptable mismatch rates in alignments or allowable run-time variance. | GSSO:009002 [The allowable deviation from a standard; a permissible difference] | http://purl.obolibrary.org/obo/GSSO_009002
Trueness | Closeness of agreement between the average of a large number of test results and a reference (true) value (ISO 5725). It reflects systematic error, not random error. | The trueness of a qPCR assay was assessed by comparing measured concentrations of a reference standard to the certified target values. | Closeness of agreement between the average result produced by a bioinformatics pipeline and a reference (true) value (ISO 5725). Trueness reflects systematic error rather than random variation and is closely linked to bias. | The trueness of a variant calling pipeline was assessed by comparing called variants against a gold-standard dataset, such as the Genome in a Bottle reference set, to identify and correct systematic biases in the analysis.
User | Individuals or organizations that use a product, service, or system (ISO 15189 / ISO 17025). Users can include patients, clinicians, researchers, or other stakeholders relying on laboratory outputs. | Users of the laboratory include clinicians requesting RNA-seq analysis for patient cancer samples. | An individual or organization that interacts with or relies on bioinformatics products, services, or systems. Users may include researchers accessing processed datasets, clinicians receiving variant reports, or other stakeholders using pipeline outputs for decision-making. User needs inform the design, documentation, and quality control of pipelines, software, and analysis workflows to ensure outputs are reliable, reproducible, and fit for purpose. | Users of the SARS-CoV-2 analysis pipeline include virologists accessing curated variant datasets and clinicians using reports to inform treatment decisions.
User Requirement Specification (URS) | A document that defines what the user needs from a system, device, or process, expressed in functional or performance terms (per ISO/IEC and software validation guidance). The URS forms the foundation for design, verification, and validation activities. | The user requirement specification stated the pipeline must process FASTQ files and produce annotated VCFs within 48 hours. | In bioinformatics, a URS defines what analysts, clinicians, or researchers require from a computational system — such as accepted input formats, expected outputs, performance targets, or reporting needs. It ensures pipelines are designed and validated to meet end-user expectations. | The user requirement specification stated that the LIMS must track sample barcodes, record reagent batches, and generate audit reports accessible within 24 hours.
Validation | Confirmation, through objective evidence, that requirements for a specific intended use are fulfilled (ISO 9000, VIM). Validation ensures that a method, assay, or procedure produces accurate, reliable, and reproducible results for its intended purpose. | Validation of the RNA-seq pipeline confirmed it was suitable for accurately quantifying gene expression in FFPE cancer samples, supporting clinical decision-making by identifying the most appropriate therapy for individual patients. | In bioinformatics, validation demonstrates that a pipeline, tool, or dataset is fit for purpose. It goes beyond technical correctness to confirm suitability for the stated analysis. | A variant calling pipeline was validated by demonstrating sensitivity and specificity on gold-standard datasets, ensuring it is appropriate for clinical use. | NCIT:C16237 [The act of validating; finding or testing the truth of something] ¦ OBI:0302911 [a planned process with objective to check that the accuracy or the quality of a claim or prediction satisfies some criteria and which is assessed by comparing with independent results] | http://purl.obolibrary.org/obo/NCIT_C16237 ¦ http://purl.obolibrary.org/obo/OBI_0302911
Verification (Verification Plan, Schedule of Verification) | Confirmation, through evidence, that a method, pipeline, or tool has been implemented correctly according to specifications (ISO 9000). Verification focuses on technical correctness and adherence to SOPs or design, rather than suitability for the broader intended purpose. | The verification plan required re-running a standard dataset quarterly to ensure the assay consistently produced expected results. | Verification of the updated SARS-CoV-2 pipeline confirmed that the new version still achieved its intended aims, producing expected results, maintaining sensitivity and specificity, and performing consistently on standard and edge-case datasets before deployment. |  | NCIT:C45513 [The act of reviewing, inspecting, testing, checking, auditing, or otherwise establishing and documenting whether items, processes, services, or documents conform to specified requirements. Compare to validation where suitability to purpose is also established] | http://purl.obolibrary.org/obo/NCIT_C45513
Workbench | A designated workspace where specific tasks are performed in a controlled way. In ISO/QMS, a workbench is physical or digital, defined by its tools and controls. | The RNA extraction workbench was organized and equipped to ensure consistent sample handling and minimize contamination. | : A workbench can be a software platform (e.g., Galaxy, Chipster) or a compute environment/cluster where analyses are performed in a controlled and reproducible manner. | The Galaxy workbench and HPC cluster were validated as controlled environments for reproducible SARS-CoV-2 data analysis.
Note on C14: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Non bioinformatics example
Note on A51: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: also in row 41
Note on A52: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: also in row 41